Supplementary Information:  Optimization of hyphenated asymmetric flow field-flow fractionation for the analysis of silver nanoparticles in aqueous solutions

Felix Geißler1, María Martínez-Cabanas1*, Pablo Lodeiro1,2 and Eric P. Achterberg1

1	Chemical Oceanography, Marine Biogeochemistry, GEOMAR Helmholtz Centre for Ocean Research Kiel, Kiel, Germany

2	Department of Chemistry, University of Lleida – AGROTECNIO-CERCA Center, Rovira Roure 191, 25198, Lleida, Spain

*	Correspondence: M.M.C.: mmartinez@geomar.de

# 1. Example of an AF4 run program

Table S 1. Cross-flow profile for the separation of AgNPs using AF4.

| No. | Step | Duration [min] | Vx [mL/min] | Remarks |
| --- | --- | --- | --- | --- |
| 1 | Elution | 1 | 0.71 | Pre-conditioning |
| 2 | Focus | 1 | 2.0 |  |
| 3 | Focus + Inject | 1 or 2 | 2.0 | Handling of AgNPs |
| 4 | Focus | 101 | 2.0 |  |
| 5 | Elution | 201 | 0.71 |  |
| 6 | Elution + Inject | 1 | 0.0 | Flushing |

1 Throughout the study presented herein, these parameters were subject of modification in order to determine the optimum conditions for the separation of AgNPs.

The purpose of the pre-conditioning steps 1 and 2 was to provide a baseline and to stabilize the pressure inside the system. In step 3 to 5 the AgNPs were injected onto the separation membrane, focused and separated according to their individual diffusion coefficients. For the injection of the sample (step 3) a duration of one minute for the 20 µL and two minutes for the 100 µL injection loop was chosen in order to ensure a complete release of the whole sample into to separation channel. The acquisition of the fractograms started with the elution step 5. The system was flushed without any cross-flow in step 6 to remove any remaining AgNPs from the separation system, observable as a ‘release peak’ in the fractograms.

# 2. UV/Vis spectra

| (a) 20 nm AgNP_LA | (b) 50 nm AgNP_LA |
| --- | --- |
| (c) 80 nm AgNP_LA | (d) mixture of 20 nm and 80 nm AgNP_LA |

Fig. S1 UV/Vis spectra of AgNPs of different diameters coated with lipoic acid in three different liquid media.

3. Fractograms of 20 nm, 50 nm and 80 nm AgNP_LA

| (a) water pH 8 | (b) 0.05% v/v Mucasol |
| --- | --- |
| (c) 1 mM NaCl |  |

Fig. S2 Fractograms of 20 nm AgNP_LA for different applied cross-flows and different carrier solutions.

| (a) water pH 8 | (b) 0.05% v/v Mucasol |
| --- | --- |
| (c) 1 mM NaCl |  |

Fig. S3 Fractograms of 50 nm AgNP_LA for different applied cross-flows and different carrier solutions.

| (a) water pH 8 | (b) 0.05% v/v Mucasol |
| --- | --- |
| (c) 1 mM NaCl |  |

Fig. S4 Fractograms of 80 nm AgNP_LA for different applied cross-flows and different carrier solutions.

# 4. Retention times for 50 nm and 80 nm AgNP_LA

| (a) 50 nm AgNP_LA | (b) 80 nm AgNP_LA |
| --- | --- |

Fig. S5 Effect of carrier solution and applied cross-flow Vx on retention time tR.

5. Aggregation of AgNP_LA in fjord water

| (a) 20 nm AgNP_LA | (b) 50 nm AgNP_LA |
| --- | --- |
| (c) 80 nm AgNP_LA | (d) temporal trend of PdI |

Fig. S6 (a, b, c) UV/Vis spectra of all three AgNP_LA size fractions acquired at 0, 10, and 300 minutes after spiking into fjord water and (d) temporal trend of polydispersity index of AgNP_LA after spiking into fjord water.

# 6. Size calibration

Based on the previous findings, AF4 can be a powerful tool to evaluate particles’ hydrodynamic diameters (dH) of a sample with unknown particles size based on their individual tR. For this purpose, size calibration (tR vs. dH) has to be conducted for each individual type of nanoparticles, carrier solutions, membranes and applied cross-flows. Fig. S7 shows size calibrations for AgNP_LA at different cross-flows and carrier solutions.

Fig. S7 Size calibration of the used AF4 system for two different cross-flows (Vx = 0.2 mL/min: closed symbols; Vx = 0.5 mL/min: open symbols) and different carrier solutions. Hydrodynamic diameters were taken from the DLS characterization (Table 2).

Table S2. Linear fits of tR vs dH plots from Fig. S7.

| Carrier solution | Vx | Linear fit | R² |
| --- | --- | --- | --- |
| Water pH 8 | 0.2 mL/min 0.5 mL/min | tR = 0.123 min∙nm-1 dH + 1.309 min tR = 0.220 min∙nm-1 dH + 1.215 min | 0.990 0.963 |
| 0.05% v/v Mucasol | 0.2 mL/min 0.5 mL/min | tR = 0.097 min∙nm-1 dH + 3.649 min tR = 0.225 min∙nm-1 dH + 2.939 min | 0.990 0.995 |
| 1 mM NaCl | 0.5 mL/min | tR = 0.624 min∙nm-1 dH – 6.915 min | 0.983 |

The slope of the obtained linear fits increased with a larger applied cross-flow (Table S2 and solid lines vs. dashed lines in Fig. S7), which implies a better separation for samples containing different sized nanoparticles. The biggest gradient, so the maximum delayed elution time between different size fractions, was achieved with the use of 1 mM NaCl as carrier solution. However, a negative intercept was obtained for this calibration. One can assume that this is the result of a deviation from the linear behavior due to enhanced particle-membrane interactions at high cross flows and 1 mM NaCl carrier solution as the size of the AgNP_LA increased. Deviations from linearity were also obtained in other studies (1,2). Moreover, the use of 1 mM NaCl has a negative effect on the duration of the separation experiment, with the elution of 80 nm AgNP_LA after 55 minutes versus 25 minutes when 0.05% v/v Mucasol was used. The anomalously long tR when using 1 mM NaCl as carrier solution, together with the low recoveries and the obtained AF4 fractogram (Fig. S5) showing a very sharp peak, indicate that the NPs were aggregated along the AF4 channel and then partially removed from the system as a release peak. Thus we selected the carrier solutions water pH 8 and 0.05% v/v Mucasol for the calibration approach.

References

1. 	Geiss O, Cascio C, Gilliland D, Franchini F, Barrero-Moreno J. Size and mass determination of silver nanoparticles in an aqueous matrix using asymmetric flow field flow fractionation coupled to inductively coupled plasma mass spectrometer and ultraviolet-visible detectors. J Chromatogr A. 2013;1321:100–8.

2. 	Gigault J, Pettibone JM, Schmitt C, Hackley VA. Rational strategy for characterization of nanoscale particles by asymmetric-flow field flow fractionation: A tutorial. Anal Chim Acta. 2014 Jan;809:9–24.